## Slide 1
© 2023 HORIBA, Ltd. All rights reserved.
Jeffrey Bodycomb, Ph.D.
Introduction to Dynamic Light Scattering (DLS)
March 23, 2023
HORIBA Scientific
Particle Characterization

## Slide 2: Outline
Introduction
What is DLS and what does it measure?
Method development

## Slide 3: What is dynamic light scattering?
Dynamic light scattering refers to measurement and interpretation of light scattering data on a microsecond time scale.
Dynamic light scattering can be used to determine
Particle/molecular size Size distribution
Relaxations in complex fluids

## Slide 4: Other light scattering techniques
Static Light Scattering: over a duration of ~1 second. Used for determining particle size (diameters greater than 10 nm), polymer molecular weight, 2nd virial coefficient, Rg.
Electrophoretic Light Scattering: use Doppler shift in scattered light to probe motion of particles due to an applied electric field. Used for determining electrophoretic mobility, zeta potential.
Nanoparticle Tracking Analysis (NTA): use scattering to track particle location as a function of time, that is, particle motion. Use motion to determine particle size.

## Slide 5: Sizing techniques
ViewSizer 3000
PSA-300
LA-350
LA-960
SZ-100

## Slide 6: Laser diffraction
Laser Diffraction
Particle size 0.01 – 5000 µm
Converts scattered light to particle size distribution
Quick, repeatable
Most common technique
Suspensions & powders

## Slide 7: Laser diffraction
q(%)
Diameter(µm)
14
2
4
6
8
10
12
0
10.00
100.0
1000
0
3000
10
30
20
40
50
70
60
80
100
90
Suspension
Silica
~ 30 nm
Powders
Coffee Results
0.3 – 1 mm

## Slide 8: Nanoparticle tracking analysis (NTA)
Video
Camera
(CCD)
Objective Lens
Scattered
light
Laser 445 nm (Blue)
Laser 520 nm (Green)
Laser 635 nm (Red)

## Slide 9: Nanoparticle tracking analysis (NTA)
microscope + camera
light sheet
scattered light
light sheet thickness
investigated volume
light source

## Slide 10: Outline
10
Introduction
What is DLS and what does it measure?
Method development

## Slide 11: DLS optics
Particles
Backscatter (173°)
(High concentration)
90°for size and MW, A2
Laser
532nm, 10mW
Attenuator
Particles moving due to Brownian motion

## Slide 12: Brownian motion
Particles in suspension undergo
Brownian motion (random thermal motion).
Brownian Motion
  Random
  Related to Size
  Related to viscosity
  Related to temperature

## Slide 13: DLS signal
Random motion of particles leads to random fluctuations in signal (due to changing constructive/destructive interference of scattered light).

## Slide 14
Correlation function
Random fluctuations are interpreted in terms of the autocorrelation function (ACF), C().
T
 I (t)I (t  )dt
C( )  1  exp(2 )
I (t)I (t)
C( )  0

## Slide 15: q  4n sin  
Gamma to size
2

m
kBT
3(T )D
Dh 
m
  D q2
 decay constant
Dm diffusion coefficient q scattering vector
n refractive index
 wavelength
 scattering angle
Dh hydrodynamic diameter
 viscosity
kB Boltzman’s constant
Note effect of temperature!

## Slide 16: What is hydrodynamic size?
DLS gives the diameter of a sphere that moves (diffuses) the same way as your sample.
Dh
Dh
Dh

## Slide 17: Hydrodynamic size
SEM (above) and TEM (below) images for NIST RM 8011
Gold Colloids
Technique | Size nm
Atomic Force Microscopy | 8.5 + 0.3
Scanning Electron Microscopy | 9.9 + 0.1
Transmission Electron Microscopy | 8.9 + 0.1
Dynamic Light Scattering | 13.5 + 0.1

## Slide 18: Nanogold data
 | Z-average Diameter, nm
Run 1 | 50.5
Run 2 | 51.1
Run 3 | 49.2
Run 4 | 51.5
Run 5 | 49.7
Run 6 | 50.9
Avg. | 50.5
St. Dev. | 0.9
COV | 1.7 %

## Slide 19: Nanogold data
 | Z-average Diameter, nm
Run 1 | 10.5
Run 2 | 10.6
Run 3 | 10.2
Run 4 | 10.5
Run 5 | 10.3
Avg. | 10.4
St. Dev. | 0.2
COV | 1.9 %
 | Z-average Diameter, nm
Avg. | 50.5
St. Dev. | 0.9
COV | 1.7 %

## Slide 20: Lab to lab comparison
Colloidal Silica
 | Mean determined Z-average size (nm) | COV (%)
Dynamic Light Scattering with SZ-100, laboratory 1 | 34.4 | 0.7
Dynamic Light Scattering with SZ-100, laboratory 2 | 34.6 | 0.3

## Slide 21: Polystyrene latex

## Slide 22: Mixtures of particles
Sum the autocorrelation functions

## Slide 23: Polydisperse samples: cumulants
For a mixture of sizes, the autocorrelation function can be interpreted in terms of cumulants. This is the most robust method of analyzing DLS data.
C( )  1  exp(2 )
m
kBT
z,h
3(T )D
D 



 


 


2
2!
 2

C( )  1  exp 2  
m
  D q2
2

Polydispersity  2
“z-average size”

## Slide 24: Z-average
Size determined from intensity weight diffusion coefficient ~1/D
Intensity weighted harmonic mean size
Dz = z-average
Si = total scattering from all of species i Di = Diameter of species all
As size goes up, so does Dz.

## Slide 25: SiO2
Run | Z-average Diameter (nm) | Polydispersity Index
1 | 473.2 | 0.127
2 | 479.5 | 0.066
3 | 478.8 | 0.077
4 | 487.7 | 0.039
Avg. | 479.8 | 0.077

## Slide 26: Polydisperse sample: ILT
A more general relationship can be given between the autocorrelation function and the size distribution. Let each size have a relation constant . The scattering from each population is then given by S(). Now we have an integral equation. Solving for S() gives us size distribution.
C( )  1  g (1) ( ) 2
g (1) ( )   S() exp( )d

## Slide 27: Bimodal sample
Nominal 20 nm and 500 nm latex run individually

## Slide 28: Bimodal sample
Mixed sample (in black)

## Slide 29: Colloidal Silica
Standard off-the- shelf Ludox
  Colloidal Silica
  Used to clarify beer, wine, and juice
Matches data from the LA-960 (laser diffraction)

## Slide 30: BSA
BSA- well characterized protein
DLS – Can be used to determine the aggregation state of the protein

## Slide 31: Protein aggregation
Unstabilized 10mg/ml lysozyme at pH 2
Lisa Cole and Ben Burnett at the Florida Institute of Technology
Can also be done with ViewSizer (NTA)

## Slide 32: Liposomes
Liposomes to target tumor growth
Size is critical to how the
liposome
  Encapsulates protein
  Functions within body
  Remains stable over
time
  Delivers the protein

## Slide 33: Outline
33
Introduction
What is DLS and what does it measure?
Method development

## Slide 34: Dust
Dust: large, rare particles in the sample
Generally not really part of the sample
Since they are rare cannot get good statistics

## Slide 35: Filters are your friend
Filter to remove dust. If particles are too large (D >50 nm for 0.1 m filter), at least filter diluent.
Filters available with pore sizes from 20 nm to 2m.
We can also centrifuge the sample and extract the supernatant.

## Slide 36: Suspension liquid
Choose a liquid that
  does not dissolve the particles
  prevents loose agglomerates
Add energy to break up loose agglomerates
  stirring
  ultrasound

## Slide 37: Surfactants
Enable wetting Prevent agglomeration
Common concentration: 0.01-0.1% example:
Tetrasodium pyrophosphate (TSPP) Triton X

## Slide 38: Wetting
Many dry particle samples will never form a nanoparticle suspension without significant effort.
Sprinkle particles on top of target dispersant. If the particles float on top and do not penetrate the water surface, they are not wetted. This is usually a bad sign.
If the particles break through surface and sink, they are a) wetted or b) so big that gravity is more important than surface tension. If it is case (a), you are in luck.
Particles floating on top
Plume of sinking particles

## Slide 39: Solvents
Working with aqueous systems is usually easier for many reasons.
But don’t forget to try a less polar solvent such as isopropyl alcohol.
Don’t forget that organic solvents are more difficult to handle due to fire and health hazards.

## Slide 40: Try a series of options
Make a series of suspensions and check them by eye, then measure.

## Slide 41: Estimate concentration by eye
Best range for this material
Read a newspaper through it

## Slide 42: Concentration
Plot measured size vs. concentration or dilution to determine concentration range. Note log scale.

## Slide 43: Effect of salt concentration
Note that when we suppress effect of charges by adding salt, the effect of concentration is suppressed.
Concentration effects are due to changes in particle motion, not just multiple scattering.

## Slide 44: Hints Summary
Web search
Consider solvent and surfactant
Consider ultrasound
Expect to filter
Choose largest cell you can
Optimize concentration

## Slide 45: Settling and DLS
Not all motion is Brownian motion 
The Natural limit for Dynamic Light Scattering: Gravitational Settling
Gravitational Settling occurs at about 1-3m
Particle Diameter (m) | Movement due to Brownian Motion |  | Movement due to Gravitational Settling
0.01 | 2.36 | >> | 0.005
0.25 | 1.49 | > | 0.0346
0.50 | 1.052 | > | 0.1384
1.0 | 0.745 | ~ | 0.554
2.5 | 0.334 | < | 13.84
10.0 | 0.236 | << | 55.4

## Slide 46
DLS disadvantages
Sensitive to large particles
Poor resolution of distribution
Not appropriate where settling is significant (use laser diffraction)

## Slide 47: DLS Advantages
Fast
Repeatable
Noninvasive
Requires only small quantities of sample
Good for detecting trace amounts of aggregate
Good for macromolecular sizing
Reaches smallest particle sizes

## Slide 48: SZ-100
Single compact unit that performs size, zeta potential, and molecular weight measurements.

## Slide 49: Summary
Fast, repeatable nanoparticle sizing
Think about suspension chemistry in method development
Reports hydrodynamic size:
Dh
Dh
Dh

## Slide 50
© 2022 HORIBA, Ltd. All rights reserved.
50